GC×IMS Compound Matching

How drift_relative is determined, and how a peak is matched to a compound from drift-rel + RI

Source modules: rip.py, match.py  ·  GC-IMS-PEAK Identify workflow, Stage 1 + Stage 5

The key idea: both sides compare the same dimensionless ratio to the Reactant Ion Peak (RIP) — never an absolute drift time. That is what lets IMS matching run without any K0 calibration.

1.  Determining the peak’s drift_relative

Produced in Stage 1 (rip.py). First locate the RIP on the drift axis: take the RT = 0 row of the intensity matrix, skip the first 200 samples, and take the position of the maximum.

r = argmaxj ≥ 200 I[0, j]

Then each peak’s own drift index jₚ (its dt_index) is divided by that RIP index to give a relative value:

dp = jpr

The RIP itself is exactly 1.0; a value dₚ = 1.13 means the peak drifted 13% farther than the RIP. Being a pure ratio, it is immune to sample rate, tube length, temperature and pressure.

2.  The library data — and the unit filter

Each .iml library row stores its drift value in Dt[a.u.], but the unit convention varies and is tagged by DtMode. The same compound can appear under different conventions:

| DtMode | What Dt[a.u.] means | Comparable to dₚ ? |
| --- | --- | --- |
| RIPrel | Ratio to the RIP (library’s own jℓ / rℓ) | Yes — same dimensionless ratio |
| 1/K0 | Inverse reduced mobility | No — needs K0 calibration |
| ms | Raw drift time (milliseconds) | No — different unit |

So match_drift_rel()’s first gate keeps only DtMode == "RIPrel" rows — the ones in the same unit as the peak. Everything else is skipped; mixing units would produce meaningless matches.

3.  The drift-axis match condition

For each surviving RIPrel row ℓ, compute the absolute difference between the library drift ratio dℓ and the peak’s dₚ, and test it against the tolerance half-width τ_d:

Δdℓ = dℓ − dp  ≤  τd = 0.05

Hits are sorted ascending by Δ_d; the closest one fills the IMS column, e.g. 1.139 (Δ0.026).

4.  Two axes intersect → the confident identification

Drift agreement alone only says “something drifts nearby.” The identification comes from intersecting with the GC/RI axis, which uses the same shape of test:

ΔRIℓ = RIℓ − RIp  ≤  τRI = 5

combine_by_cas() then keys on the CAS number and keeps only compounds whose CAS lands inside both windows:

combined = c : ΔRI ≤ τRI ,  CAS = c  ∧  Δd ≤ τd ,  CAS = c }

The intersection is ranked by the combined error Δ_RI + Δ_d and shown in the GC×IMS column. Agreement on both axes collapses the RI axis’s hundreds of single-axis candidates down to a few.

5.  End-to-end flow (match_all for one peak)

Amber = GC/RI axis · Teal = IMS/drift axis · Indigo = both. The K0 branch is currently dormant.

Current status

The K0 branch is dormant (no calibration standard yet), so the IMS axis runs on match_drift_rel (RIPrel). Both windows, τ_d = 0.05 and τ_RI = 5, are still uncalibrated placeholders and should be validated against known compounds on a reference run.